## Slide 1: High Threshold Cerenkov Counter (HTCC) For CLAS12
Hakob Voskanyan
(AANL, Armenia)

## Slide 2: Outline
• General Description
Parameters
Main Components
• Light Monitoring System
Components
Calibration

## Slide 3: Clas12 Detector

## Slide 4: HTCC and Central Detectors

## Slide 5: High Threshold Cerenkov Counter (HTCC) for CLAC12
HTCC

## Slide 6: HTCC Parameters
PARAMETER DESIGN | DESIGN VALUE
Working Gas | @1atm, 25⁰C
Angular Coverage | ϑ = 5⁰ – 35⁰; ϕ = 0⁰ – 360⁰
Threshold | 15 MeV/c (electrons)
Threshold | 4.9 GeV/c (charged pions)
Rejection of pions at 2 GeV/c | (99.9% electron detection efficiency)
Rejection of pions at 4 GeV/c | (99.9% electron detection efficiency)
Overall Dimensions | 15 ft and L = 6 ft along beam direction
Mirror Type | Combined Ellipsoidal, composite, 8 ft in diameter
Mirror Substrate Structure | Acryl (0.01”) + PMI Foam (0.6”) + Acryl (0.01”)
Mirror Thickness | (130 – 135) mg/cm ²
Number of Mirror Half Sectors | 12
Number of Mirror Segments | (12x5) = 60
Number of Channels | (12x4) = 48
Photomultiplier Tubes | Electron Tubes 9823QKB (5”, quartz window)
Number of Reflections | 1 (80%) + 2 (20%)
Typical Magnetic Field | 35 Gauss (along PMT axis)
Standard Magnetic Shield | Multilayer with compensation coil
PARAMETER DESIGN | DESIGN VALUE
Working Gas | 
Angular Coverage | ϑ = 5⁰ – 35⁰; ϕ = 0⁰ – 360⁰
Threshold | 15 MeV/c (electrons)
Threshold | 4.9 GeV/c (charged pions)
Rejection of pions at 2 GeV/c | 
Rejection of pions at 4 GeV/c | 
Overall Dimensions | 15 ft and L = 6 ft along beam direction
Mirror Type | Combined Ellipsoidal, composite, 8 ft in diameter
Mirror Substrate Structure | Acryl (0.01”) + PMI Foam (0.6”) + Acryl (0.01”)
Mirror Thickness | (130 – 135) mg/cm ²
Number of Mirror Half Sectors | 12
Number of Mirror Segments | (12x5) = 60
Number of Channels | (12x4) = 48
Photomultiplier Tubes | Electron Tubes 9823QKB (5”, quartz window)
Number of Reflections | 1 (80%) + 2 (20%)
Typical Magnetic Field | 35 Gauss (along PMT axis)
Standard Magnetic Shield | Multilayer with compensation coil

## Slide 7: HTCC’s Main Components
1. COMPOSITE ENTRY WINDOW
2. ENTRY DISH
3. ENTRY CONE
4. MOLLER CUP
5. PMT HOUSING UNITS
PMTs and Winston Cones
Magnetic Shields
Compensation Coils
Modified HV-dividers
6. WELDEMENT PART
7. OUTER SHELLS
8. COMBINED MIRROR
9. EXIT WINDOW

## Slide 8: High Threshold Cerenkov Counter (HTCC) for CLAC12
Main Components

## Slide 9: PMT HOUSING UNITS
PMT HOUSING UNITS
PMTs and Winston Cones
Magnetic Shields
Compensation Coils
Modified HV-dividers

## Slide 10: Installation of PMT Mounting Unit
Assembled Unit
Installation of first four Units
Installed Units

## Slide 11: HTCC Combined Ellipsoidal Mirrors
Combined Ellipsoidal Mirror
Each sector of the CLAS12 is covered with 2 identical Half-sector Mirrors.
The HTCC is a one unit detector with core component of Ellipsoidal Combined Mirror (consisting of 60 lightweight Ellipsoidal Mirror Facets).
Mirror Thickness - (130 – 135) mg/cm²

## Slide 12: Geometry
The First focal point always corresponds to the target location, and the top focal point always corresponds to the PMT location

## Slide 13: Mirrors construction
Tooling for mirrors construction
Substrate M3 ready for coating.
Front is covered with KAPTON
tape to prevent deposition of
coating material

## Slide 14: Coating and Reflectance Tests
o Free of visible defects
o Test results acceptable and consistent with flat
sample results used to award contract to ECI

## Slide 15: Assembly of the HTCC Half-sector Mirrors
- Epoxy glue of high viscosity and small
shrinkage was applied to one side only
covering ~25% of contact area.
- Geometry was tested before and after
facets were glued together.
- Environment in the HTCC Clean
Room was RH=36%, T=79.8⁰F.
- Assembly steps are illustrated in
pictures bellow.
#4
#3
#2
#1Right
#1Left
High Precision Base for single
Half Sector assembly

## Slide 16: Geometry checked with Laser beam rasterized off the center line of Half-sector Mirror
Adjustable Laser mount
Focal Planes

## Slide 17: Combined Mirror Assembly
Combined Mirror Assembly Jig
Combined Mirror Assembly Jig allows:
Step-by-step assembly of the Combined Mirror by adding
one Half Sector Mirror at the time and start gluing
Process after all 12 Half Sectors aligned
Vacuum retention of the Half Sectors (1 to 12 in any
combination) for extended period of time
Combined Mirror Assembly Jig
OD=8ft.

## Slide 18: Mirror installed in HTCC

## Slide 19: PMT Alignment

## Slide 20: HTCC Light Monitor System (LMS)
(1) Integrating Sphere
(2) Light Source
(3) Adapters
(4) Reference PMT with and YAP
(5) Harness and Fiber Bundle
(6) PIN diode
(7) Magnetic Shield
1
2
3
4
5
6
7
Fast Checking the detector operation
Monitoring gain of all HTCC channels
Calibration each channels

## Slide 21: HTCC Light Monitor System

## Slide 22: Test Setup
PMT base (HV Divider + Amplifier)
PMT Gain Matching
Light source (LED + driver)
Fiber Bundle
Test purpose
Dark Box

## Slide 23: Test Setup
PMT
fADC250
CODA
Function Generator
HV Power Supply
LED
Integrating Sphere
Discriminator
LED Driver

## Slide 24: Test Setup
PMT
LED
Integrating Sphere
LED Driver

## Slide 25: Integrating Sphere
Integrating Sphere

## Slide 26: Test Setup
Winston Cone with Fiber

## Slide 27: Test Setup
Electronics Setup

## Slide 28: PMT base (HV Divider + Amplifier)
ET base, external amplifier HV=2.2kV
Modified active base at HV=1.6kV
No positive bump on the tail
Leading edge 2.6ns 3.8ns
FWHM 5.4ns 8.0ns
Total width 10ns 20ns
Output range 2.5V 3V on each output
Outputs number 1 2

## Slide 29: PMT Gain Matching
48 PMTs are used in the fast trigger and must be gain
equalized
The face of the PMT will be impinged with an LED pulse
provided by the fibers during a special calibration run
The ADC spectra will be fit to find the position of the single
photo electron (SPE) peak
The positions of the SPE peak and correspond calibration
curve (SPE peak vs ) for each PMT will allow for gain
matching of the PMTs

## Slide 30: A Model of PMT response
Photo conversion and
electron collection
Amplification
P(n; μ ) = = mq is a mean number of photoelectrons collected by the first dynode
(x) =
(x) =
Amplification
= P(n; )=
Operation for a PMT

## Slide 31: A Model of PMT response
Background processes
The low charge processes:
the leakage current, external and internal radioactivity, etc.
The discrete processes :
thermoemission, noise initiated
by the measured light, etc.
B(x)
= (x-
erf()]

## Slide 32: Fit Results

## Slide 33: Pavel Fit

## Slide 34: HV test
Fitting the position of the SPE peak as function of high voltage we can equalize all channels

## Slide 35: Intensity Test
Shows the range of µ, and what range the LED driver can cover
This test will be repeated with µ ranging from .3-1.0
PMT HV =1600V

## Slide 36: Intensity Test
• Gain of the PMT held
constant
• Good results with
Varying light intensities
Fit results compared
With other algorithms

## Slide 37: Finding the Calibration Curve
• Gain of the PMT
measured with fit
• Each PMT will have it’s
own gain curve
• Fit results also
compare with other
algorithms

## Slide 38: Fiber Bundle Test
All 50 fibers produced a signals
Intensity is uniform ± 3%
μ